XG Distribution Disk README

This file contains last minute release information

Version 3.01			Date: November 24, 2014

- 485 interface was not working in Ethernet mode.  Added an option from the front panel for selecting “485L”.  Using this mode will allow 485 multidrop to work in LAN mode.

- Added fix to prevent buffer overrun when mutlitple carriage returns and line feeds are transmitted.

- Added fix to prevent voltage and current setting from jumping when voltage/current are being adjusting using the rotary knob at a fast rate.

- Changed update to use manufacturer provided bootloader instead of IAP code.

- NXP processor is EOL.  Code needs to change to support SST replacement processor.

Version 2.29		Date: April 4, 2014

- Fixed Interlock behavior not allowing the OUTP:PROT:CLE to clear alarm while interlock is enabled.

- Adds handshake mechanism with Ethernet communication.

Version 2.27		Date: November 22, 2013

- Added handshaking functionality to the Ethernet communication side only.

Version 2.25		Date: August 14, 2013

- Modified code to automatically fix calibration data if bad calibration data is detected.

Version 2.23		Date: March 8, 2013

- Masking External Shutdown is performed through SCPI commands.  Refer to page 3-22 in the XG User Manual for more information about Alarm Masking.  The command to mask only the External Shutdown is "SYST:PROT:MASK 1535" the default command for is "SYST:PROT:MASK 2047".  The tables below compare the difference when External Shutdown mask is ON/OFF in the current firmware 2.21 and the new firmware 2.23.

- XG Flash Updater now allows upgradability from 2.21 to 2.23 and from 2.23 to future firmware.

- Seabotix issue with voltage dropping is now resolved in Firmware 2.23.

Version 2.21		Date: October 17, 2012

- After setting foldback to "CV" mode and rechecking which protection mode its in the front panel will display "FLD NONE" when it should display "FLD CV".  This firmware corrects this problem.

- New SCPI command "DIAG:FINE?/OFF/ON" added to allow "FINE" tune voltage and current adjustment instead of the default COARSE.

- A new feature has been added Constant Power Mode.

- The Loop Protection is used to enable or disable multiply units via daisy chain configuration through the AUX Output and Isolated