# Multioutput Front Panel Firmware Release Notes

V1.113

- Fixed issues related System Status menu.

V1.110

- Implemented  Self-Test features.

V1.109

- Fix for Remote update issue

V1.108

- Added LXI logo in the start-up screen.

V1.107

- Changed the order of Unit Input Voltage (VAC) order in *idn?

V1.106

- Removed EtherCAT from Control Interface menu.

V1.105

- Unable to enable the output from front panel issue is fixed.

- In User VI limits for data not stable when set using encoder issue is fixed.

- Removed LXI logo from the start up screen.

- Data not stable in PONS Menu issue is fixed.

- Lockout screen implemented with password.

- Rotary Encoder issues are fixed.

V1.104

- Unable to enable the output from front panel issue is fixed.

- User VI limits for data not stable when set using encoder issue is fixed.

V1.103

- Have fixed for the bugs mentioned below.

  - When remote interface is initialized and if unlock button is pressed power supply should allow access to the user to control the power supply through the front panel also,

  - if command issued from the remote interface the screen should be locked again.

  - Output on/off delay --> configure delay --> unselected channel's delay setting should be masked.

  - Output on/off delay --> configure delay --> channel 1 is always masked and selected ref channel delay settings should be masked.

  - Output ON/OFF delay --> Settings --> programmed delay are not updated in the settings screen

  - Output ON/OFF Delay --> after some time unable to select reference channel, unable to set delay.

  - Configuration --> unable to scroll configuration menu screen 2

  - PONS Analog --> unable to scroll  menu screen 2

  - PONS - Current reference mode --> unable to select mode, it keep switching between options.

  - PONS – Reference source settings --> current ref source is not displayed and current ref FSC is displayed as NA, while reference mode is in external.

  - Reference source settings --> current ref source is not displayed and current ref FSC is displayed as NA, while reference mode is in external.

  - Output on/off delay --> unable to initiate the output sequence from the front panel.

- Fixed Encoder issues in mentioned below screens.

  - Measurement settings

  - User VI Limits

  - PONS - Measurement settings

  - PONS - User VI Limits settings

  - PONS - Output Program settings

  - Output Program - Voltage Settings

  - Output Program - Current Settings

  - Output Program - OVP Settings

  - Output Program - Power Settings

  - Regulation Mode - Delay Settings

V1.102

- 1.Implementation of Output On/Off Delay Menu as per new requirement

- 2.Implementation of Multichassis feature as per new requirement.

- 3.Implementationof Analog menu as per new requirement.

- Have fixed for the bugs mentioned below.

- As of now we are not going to release multi chassis feature, the standalone screen should not appear.

- When remote interface is initialized and if unlock button is pressed power supply should allow access to the user to control the power supply through the front panel also,

  - if command issued from the remote interface the screen should be locked again.

- Remote inhibit--> live selected from the front panel while querying using SCPI command it is in latch.

- 4.	Remote inhibit--> if any one of the channel is selected as live or latch the other channels also selected automatically the same.

- 5.	Output on/off delay --> configure delay --> unselected channel's delay setting should be masked.

- 6.	Output on/off delay --> configure delay --> channel 1 is always masked and selected ref channel delay settings should be masked.

- 7.	PONS - Current reference mode --> unable to select mode, it keep switching between options.

- 8.	PONS – Reference source settings --> current ref source is not displayed and current ref FSC is displayed as NA, while reference mode is in external.

- 9.	Reference source settings --> current ref source is not displayed and current ref FSC is displayed as NA, while reference mode is in external.

- 10.	disabled settings field box in dashboard screen allows to program using rotary encoder.

- 11.	disabled settings field box in dashboard screen allows to program using touch screen.

- 12.	Output on/off delay --> unable to initiate the output sequence from the front panel.

V1.098

- 1.Implementation of Output On Delay Menu

- 2.Implementation of System Status Menu for all 3 channels

- 3.Implementation of Factory Default Settings Menu

- 4.Implementation of Multichassis feature.

- 5.Enhancement of Analog screen full scale values.

- 6.Issues reported in Bugzilla is fixed.

  - 917 Configuration menu - PONS - Analog - Configure Analog

  - 919 Regulation mode status is not getting updated in the measurement’s screen/default screen

  - 924 On UUT Runtime, while setting Configuration --> User V/I limits, the values entered are being reflected under PONS --> Output Program (minimum and maximum values).

  - 927 Enhancement of Analog screen

V1.095

- LAN configuration Menu not highlighted issues fixed.

- when UTT is booted up with PONS output enable, output warning screen is popped up. When Abort is pressed Aborting message will be displayed.

- Issue reported in Bugzilla fixed:

- 926 Foldback Fault screen

- 925 E-stop Fault conditions

- 923 Control Interface - LAN - Configure LAN

V1.094

- Factory Default screen added

- Analog FSC/FSCR value setting issues fixed

- LAN Static IP Configuration issues fixed

V1.0.93

- GPIB issue is fixed

- LAN Configuration issues fixed

- CP label issue in Dashboard menu and Measurement settings fixed

V1.0.91

- Implemented PONS Output enabled warning screen :

- BugZilla Issues Fix:

- 897-Power-up screen

- 900-Control Interface -RS232

- LAN AUTO IP, PORT No, Host name not setting issue fixed

V1.0.9

- Issue reported in BugZilla fixed:

- 877	Dashboard screen Output ON/OFF

- 878	Voltage and current setting in Dashboard screen

- 879	Output program --Output state

- 880	Output-Program-Regulation modes

- 887	Not able to clear the error screen from front panel

- 884	Control interface--RS232

- 888	Front panel download message jumbled

- 890	Dashboard screen - voltage and current settings

- 891	Output Program menu - Power settings

- 893	Configuration menu - User VI settings

- 894	PONS ->User VI settings

V1.0.89

- Dashboard Screen implemented

- Output Menu: Output program, Output State, Regulation mode and Output sense screen implemented

- System settings menu: Hardware limits, LCD, Firmware version and default screen