# SGX Master Firmware

# Supported Instruments:

# SGX

# Release Notes:

V1.67:

- Bug Fix: Since there was build issue, in few K70 Som’s we could see failure to mount NAND, which caused font file issue. This is fixed in the firmware version and tested in multiple SOM boards.

V1.64:

- Add capability for the user to set and change their own password to lockout the display.  Need to define min/max character limits.

- Add an AMETEK master password that only Tech Support knows to override any user defined password.  Useful in case the customer forgets their password, or the unit comes to the repair centre on an RMA.

- Add capability for the user to set and change their own password to login to LXI webpage.

- Once LXI reset is done, LXI webpage login details will also to be reset to user id: admin and password: password.

### V1.60:			Date: 27-July-2023

### •	Added support for parallel up to 15 units.

### V1.57:			Date: 07-Dec-2021

- Fixed Sequencing related issues by handling the RETURN subcall.

### V1.56:			Date: 4-March-2021

- IP Reset button made to work when the device is in Remote mode.

### V1.55:			Date: 1-Feb-2021

- Fixed intermittent Sequencing screen freeze up issue.

### V1.51:			Date: 28-April-2020

- Added support for new touch controller (SX8651IWLTRT)

- Changes are made to work with both old and new touch controllers.

### V1.50:			Date: 12-February-2020

- Sequence bug fixed when the sequence contains more than one SUBCALL in the steps.

### V1.49:			Date: 10-January-2020

- Corrected Sequence issue when more than one sequence is saved in the unit.

- "Please Wait" Message added during boot up of Front Panel instead of blank screen.

- Fixed issue related to External shutdown when the unit is in Regulation mode CC, shutdown through Analog control connector does not work.

- Fixed issue related to Interactive control of Webpage when wrong command is queried, interactive control stops working until reboot.

### V1.46:			Date: 8-November-2019

- Corrected Front Panel Freeze issues and Periodic Session Termination issues.

### V1.45:			Date: 18-September-2019

- GPIB Issue of getting spurious characters in the response to a sequence of SCPI commands is corrected.

### V1.44:			Date: 16-July-2019

### IP Gateway fix for static mode

### Front Panel freeze issue during the change of volt/curr using rotary knob is addressed

### Rear Panel LEDs behaviour corrected

### Front Panel Sequencing menu issues are fixed

### V1.41:			Date: 21-March-2019

- FIX: LAN Reset button not working while SGX is in "Remote Mode".

### V1.40:			Date: 12-March-2019

- Local Lockout functionality Implemented

- CV LED turned on during boot up

- Model number String length increased to 24

- Setting netmask issue is corrected.

- FIX: OVP error appearing when pressing Apply button in Lan Configuration screen

### V1.39:			Date: 31-January-2019

- FIX: Slave units configured in Parallel will now display "AUXILIARY UNIT" on the front panel display.

- FIX: Resolved issue related to current in User Software Limits, Total System Current, and Hardware Limits Screen

- The identification parsing is corrected in front panel

### V1.34:			Date: 24-October-2018

- Fixed issues for LXI 1.5 conformance.

### V1.29:			Date: 1-August-2018

- Initial Release